## Slide 1: Leonard Maaya
Candidate for the Biostatistician position at
BIOCARTIS

## Slide 2
Data issues/QC
Unnamed first column has no #72
Remarked rows (Between B and C Small spot; error repeat run)
  How should these be handled; needs discussion to or not to include in the analyses
    For this analysis, they are part of the results
    Analysis excluding the two remarked rows available at the end of the slides
Objective
To determine the limit of detection of the Idylla™ SARS-CoV-2 using a sample dilution series of positive clinical samples collected at a range of input levels (copies/ml) each with 24 cartridges.

## Slide 3
Cartridge detection (Positive call) rates vs input level
Input Level (copies/ml) | N | Detected (%)
125 | 24 | 18(0.75)
250 | 24 | 22(0.92)
500 | 24 | 24(1.00)
750 | 24 | 24(1.00)
1000 | 24 | 24(1.00)
2000 | 24 | 24(1.00)
4000 | 24 | 24(1.00)

## Slide 4
Positive call rates vs input levels

## Slide 5
Statistical model
Model: Logistic regression
To find the lowest input level; find predicted positive rates from the model for a grid of values [125, 4000]
Lowest input level (cps/ml) with 95% positive call rates:
  282
Corresponding 95% Wilson CI for the input level (cps/ml):
  c(146, 418)
Parameter | Est. | Std. Error | z value | Pr(>¦z¦)
 | -0.46 | 1.01 | -0.45 | 0.65
 | 0.01 | 0.01 | 2.11 | 0.03
Parameter | Est. | Std. Error | z value | Pr(>¦z¦)
 | -0.46 | 1.01 | -0.45 | 0.65
 | 0.01 | 0.01 | 2.11 | 0.03

## Slide 6
Possible consideration: Probit model

## Slide 7
To be checked: Operator - Date vs Positive call rates

## Slide 8
Modeling
Include more variables in the analysis?
Bayesian?
  Prior information may help avoid including input values beyond 500 cps/ml when determining LoD
Automate analysis ?
DoE?
To determine the optimal number of cartridges for a similar exercise in the future?
To determine the factors and optimal ranges for the factors for a similar exercise in the future
  Simulation study?

## Slide 9
Thank you for your time and attention.

## Slide 10
To be checked
Operator vs Positive call rates
Day vs Positive call rates
Day-Operator vs Positive call rates interaction

## Slide 11
Operator vs Positive call rates

## Slide 12
Date vs Positive call rates

## Slide 13
Cartridge detection (Positive call) rates vs input level
Input Level (copies/ml) | N | Detected (%)
125 | 24 | 18(0.75)
250 | 23 | 21(0.91)
500 | 23 | 23(1.00)
750 | 24 | 24(1.00)
1000 | 24 | 24(1.00)
2000 | 24 | 24(1.00)
4000 | 24 | 24(1.00)

## Slide 14
Positive call rates vs input levels

## Slide 15
Statistical model
Model: Binomial
  Fixed effects
Lowest input level (cps/ml) with 95% positive call rates:
  285
Corresponding 95% Wilson CI for the input level (cps/ml):
  c(147, 423)